### Detailed Technical Description of a Fourier Transform Microwave (FTMW) Spectrometer Control System

The Fourier Transform Microwave (FTMW) spectrometer is an instrument used in rotational spectroscopy for probing the molecular structure and dynamics. The control system developed for this custom-built FTMW spectrometer is a modularized and automated framework written in Python. It integrates control over several hardware components such as the Valon frequency synthesizer, Zaber linear stages, oscilloscope, and SRS signal generator, each of which is critical for conducting experiments that involve cavity tuning, frequency sweeps, and data acquisition.

This technical description details the individual modules that interface with these devices and the overall experimental workflow. Additionally, we explore how the Python-based control system enables precise cavity tuning and high-resolution data acquisition for rotational spectroscopic measurements.

### 1. System Architecture: Hardware Integration and Python Modules

The architecture of the control system is modular, with different Python modules handling specific hardware components. Each module abstracts the underlying serial or TCP/IP communication, providing a clean interface for issuing commands, setting parameters, and retrieving data. The key hardware components include:

- Valon Frequency Synthesizer: Provides precise microwave signals in the range of MHz to GHz.

- Zaber Linear Stage: Positions the sample or cavity precisely with micrometer-level accuracy.

- Oscilloscope: Captures the time-domain signal, which is processed to yield frequency-domain data.

- SRS Signal Generator: Synchronizes microwave pulse generation with data acquisition.

These devices are controlled through Python modules. Below, we dive into the specifics of each hardware component and how the Python modules facilitate their operation.

### 2. Valon Frequency Synthesizer: Frequency Control and Stepping

The Valon frequency synthesizer generates the microwave signals required for rotational spectroscopy. The Python module valonController.py manages communication with the synthesizer over a serial connection (pySerial). This module sets the synthesizer’s mode, adjusts the RF power level, and controls frequency stepping during experiments.

#### 2.1 Serial Communication with the Valon Synthesizer

The Valon synthesizer communicates over the RS-232 protocol using ASCII commands. The Python script establishes the connection using the serial.Serial() object, which is configured with parameters like baud rate (9600) and timeout (3 seconds). After the serial connection is opened, commands such as MODe CW (to set the mode to Continuous Wave) or Frequency (to set or query the frequency) are sent.

valonConnect = serial.Serial(port='COM3', baudrate=9600, timeout=3)

This communication is abstracted by the writeValonCommand() function, which formats the commands as ASCII strings terminated by a carriage return (\r) and sends them over the serial connection. The Valon responds with a message that is parsed by the program:

def writeValonCommand(cmd):

formatCmd = f"{cmd}\r"

valonConnect.write(formatCmd.encode())  # Send ASCII-encoded command

response_bytes = valonConnect.read(1024)  # Read the response from Valon

return response_bytes.decode().strip()  # Decode and strip the response

#### 2.2 Frequency Tuning and Stepping

The Valon synthesizer’s primary role is to generate precise microwave signals. The module allows users to set the total frequency (totalFreq), from which the AWG frequency is subtracted to calculate the Valon frequency (valonFreq). This allows the synthesizer to automatically adjust to the correct microwave frequency for the experiment:

valonFreq = totalFreq - awgFreq

valonController.writeValonCommand(f"Frequency {valonFreq} MHz")

For experiments that involve frequency sweeps, where multiple frequency points need to be tested, the system automatically steps the sweep settings through the frequency using the  valonSWEEPMode(Start,Stop,Step,Rate) function, which allows for the start, stop, step, and rate to be setThis is critical for scanning a range of frequencies to identify the resonant transitions in the molecular sample.

def valonSWEEPMode(Start,Stop,Step,Rate):
    valonStartSweep(Start)
    valonStopSweep(Stop)
    valonSetStepSizeSWEEP(Step)
    valonSweepRate(Rate)
    writeValonCommand("MODe SWEep")

# Each individual function can be called individually if we want to only change one,

The sweep continues until the stop frequency is reached, automating the process of scanning the entire frequency range.

### 3. Zaber Linear Stage: Sample Positioning and Cavity Tuning

The Zaber linear stage is responsible for precise cavity-mirror positioning. Accurate positioning is critical in rotational spectroscopy since the alignment of the mirror within the microwave cavity significantly impacts the measurement quality.

#### 3.1 Zaber Device Communication and Position Control

The Zaber linear stage is controlled through the zaber.serial.AsciiSerial and AsciiDevice classes, which provide an interface for sending commands and receiving status updates from the Zaber device over a serial connection (COM4).

The Zaber device operates in microsteps, where each millimeter of movement is translated into 20997.375 Zaber steps. The module abstracts this complexity by converting user inputs in millimeters into steps:

startPosZaber = round(startPosZaber * ZABER_MM_TO_POSITION)

The Zaber stage can move to both absolute and relative positions. The moveToZaber() function instructs the Zaber to move to an absolute position in microsteps, while moveByZaber() moves the stage by a relative distance. This flexibility allows for precise cavity tuning, which is essential when aligning the molecular sample within the microwave cavity.

zaberDevice.move_abs(startPosZaber)

zaberDevice.move_abs(endPosZaber)

#### 3.2 Cavity Tuning and Speed Control

Cavity tuning is achieved by moving the Zaber stage to specific positions that maximize the signal response. The system predicts the starting position based on a calibration plot (or this can be switched to a mode where the user finds the starting position manually), ensuring that the cavity or sample is positioned optimally at the start of the experiment.

startPosZaber = ((4.2349e-9 * (totalFreq ** 2)) + (0.002142 * totalFreq) - 2.0577) - 1

The Zaber stage’s speed can be controlled using the zaberSetSpeed() function, which sets the maximum velocity in steps per second. Speed control is essential when the stage must move quickly over large distances or slowly during fine adjustments for cavity tuning.

zaberController.zaberSetSpeed(speedZaber)

The mirror's alignment inside the cavity must be precise. A cavity that is not tuned can lead to weak spectroscopic measurements. Therefore, the ability to fine-tune the cavity positioning using the Zaber stage ensures optimal sensitivity and resolution.

### 4. Oscilloscope: Data Acquisition and Fourier Transform Processing

The oscilloscope captures the time-domain signals generated by the microwave pulses interacting with the molecular sample. This signal is then transformed into the frequency domain using Fourier Transform (FFT) methods to identify rotational transitions.

#### 4.1 Oscilloscope Initialization and Setup

The oscilloscopeController.py module interfaces with the oscilloscope over a TCP/IP connection, utilizing the VISA protocol. The oscilloscope is initialized using the initializeScope() function, which establishes communication and configures the instrument’s horizontal and vertical scales, as well as its sampling rate and acquisition modes.

oscilloScope = rm.open_resource('TCPIP0::169.254.23.223::inst0::INSTR')

oscilloScope.timeout = 10000

oscilloScope.write('*CLS')  # Clear previous errors or commands

The oscilloscope is further configured through a series of SCPI (Standard Commands for Programmable Instruments) commands. These commands define the timebase (HORizontal:SCALE) and voltage scale (CHANnel1:SCALe) for capturing the microwave signal:

writeOscCmd('HORizontal:MODE:SCAle 200E-9')  # Set horizontal scale

writeOscCmd('CHANnel1:SCALe 400E-6')  # Set vertical scale

#### 4.2 Fourier Transform Data Acquisition

Once the oscilloscope is configured, it captures the time-domain signal, which is then converted into the frequency domain using an FFT. This is handled by configuring the oscilloscope’s math functions to perform the spectral transformation on the acquired signal:

writeOscCmd('MATH4:DEFINE "SpectralMag(AVG(CH1))"')

The spectral data is acquired in binary format using the curve? command. The binary data is converted into a NumPy array for further processing, allowing the system to extract the spectral information from the acquired signal:

bin_wave = oscilloScope.query_binary_values('curve?', datatype='f', container=np.array, is_big_endian=True)

The acquireFFTDataAtMax() function handles the acquisition and conversion of the time-domain signal into frequency-domain data. This data is crucial for identifying the molecular rotational transitions, which appear as peaks in the frequency spectrum.

### 5. SRS Signal Generator: Pulse Synchronization and Triggering

The SRS signal generator is responsible for generating trigger pulses that synchronize the various components of the experiment.

#### 5.1 Pulse Configuration and Synchronization

The SRScontroller.py module provides synchronization triggers that ensure the microwave pulses are generated at precise intervals, the gas valve is opened, and the oscilloscope captures the data at the correct times.

SRScontroller.setTrig(5)  # Set trigger interval

SRScontroller.startPulse()  # Start the pulse sequence

The synchronized pulses ensure that the FTMW spectrometer operates cohesively, with the Valon frequency synthesizer, Zaber linear stage, and oscilloscope all working in unison. This synchronization is critical for capturing accurate rotational spectra, as the timing of the microwave pulses directly affects the signal's phase and frequency.

### 6. Experimental Workflow and Automation

#### 6.1 Automated Frequency Scanning and Cavity Tuning

The system automatically tunes the frequency using the Valon synthesizer and adjusts the cavity using the Zaber stage. The experimental parameters, including the starting frequency, step size, and stop frequency, are input by the user at the start of the experiment. The system creates directories for data storage, ensuring that each experimental run is logged systematically:

filename = input('What name to add to file? : ')

if not os.path.exists(f'{folderpath}/{filename}_{k}'):

os.makedirs(f'{folderpath}/{filename}_{k}')

The system then runs the experiment iteratively, stepping the frequency and acquiring data at each step. The data is saved in CSV format, with each frequency scan appended to the main output file.

#### 6.2 Multithreaded Data Acquisition

The code uses multithreading to manage data acquisition and Zaber stage movements concurrently. The threading.Thread() object is used to run the Zaber and oscilloscope acquisition processes in parallel:

threadZaber = threading.Thread(target=zaberThread)

threadAcquire = threading.Thread(target=acquireThread)

This approach allows the experiment to proceed efficiently, with the system acquiring data and adjusting the cavity simultaneously.

### 7. Data Storage and Visualization

The acquired data, including the Zaber stage positions and the frequency-domain signals from the oscilloscope, are saved to CSV files for later analysis. The code creates a new directory for each experiment run, ensuring that the data is organized systematically. Each frequency sweep is appended to the main output file, allowing the user to easily analyze the results.

DF1 = pd.DataFrame({'Frequency (MHz)': xxValues, 'Intensity': yyValues})

DF1.to_csv(f'{directory}/{filename}.csv', mode='a', index=True)  # Append to main file

Additionally, the system provides real-time visualization of the acquired data using matplotlib. This enables the user to see the results of the experiment as it progresses, providing immediate feedback on the quality of the data.

plt.plot(xWave, yWave)

plt.title("Zaber Position vs. Intensity")

plt.show(block=False)

### Conclusion

The Python-based control system for the FTMW spectrometer provides a powerful and flexible framework for conducting automated rotational spectroscopic measurements. Through its modular design, the system manages all critical components, including the Valon frequency synthesizer, Zaber linear stage, oscilloscope, and SRS signal generator.